## Slide 1: B18 Probton
ASIC Solution for Optical Probe Station
Discord name | Affiliation | Experience | Role
timlam0531 | HKUST, Hong Kong | Hobbyist | Team lead
abraaaar | Toronto Metropolitan University, Toronto | Undergraduate student | Team member
belkacem | University of Tlemcen, Algeria | PhD | Team member
nitin_i | University of California, Santa Barbara | AMS Design Engineer | Team member
rezasetiabekti8375 | University College Dublin, Ireland | Master student | Team member
ashmita1509 | University College Dublin, Ireland | Master student | Team member
vik_lyn_ | University of Waterloo | Undergraduate student | Team member
anna_b75_06065 | None | Hobbyist | Team member
CadMiuM#9906 | HKUST, Hong Kong | Undergraduate student | Team member
Looking for members |  |  | Mixed Signal
Review Date:
July 4, 2024 Review Version:
Schematic Review Video (not updated): https://youtu.be/nWU8KJn_Hf8

## Slide 2
Design Objective:
Extract alignment direction using voting filter while automatically isolating and discarding corrupt signal data.
Link: https://doi.org/10.1016/j.ohx.2022.e00317
Key Specifications:
  Sensitivity: 25.0 nm/V
  Resolution:
  Backlash: 160 – 195 nm
  Step error: ~ 7%
XY-Stage [1]
Photodetector
Thorlabs FDSP625 Pigtailed Si Photodiode
Link: https://www.thorlabs.com/thorproduct.cfm?partnumber=FDSP625
Key Specifications:
  Spectral Band: 320 nm to 1000 nm
  Response Speed: 700 ps rise/fall time
  Noise: 0.01 nA dark current
Link:
https://doi.org/10.1088/0957-4484/24/45/455503
Key Specifications:
  XY Range: 5 µm
  Nonlinearity: 2.11% X / 2.73% Y
  Noise / Resolution: ~0.06 nm
  Resonance: ~360 Hz (X/Y)
MEMS
Actuator [2]
Buzzer-Scanner (Wang et al. 2013)
Photodetector
Thorlabs FDSP625 Pigtailed Si Photodiode
Intended use:
Deliver low-latency direction indicators and signal intensity data to drive motorized alignment hardware.

## Slide 3
Design Specification
5MHz clock speed
Raw readout: 700ps rise fall time, 1 μA - 40 μA, 3.5pF
Allowed MEMS frequency : 2.38 Hz ~ 156.24761 kHz 16 bit Frequency Control Word 21 bit Phase Accumulator
Low Frequency Filter ~160 kHz
Mems_drv compatible to control External MEMS DAC Driver (±10V Triangular Wave)
Wave mixer and comparator able to operate at 1.65V midrail to isolate sign of dc_offset for directional identification
System is able to identify real signal from noise and jittering is minimize during idle and calibration

## Slide 4
Decision - Simplify and Reduce to core function
Based on Proposal Review Feedback
Skip unnecessary signal processing component that enhance accuracy of prediction Removed: Integrator, Voltage to Time Converter, Time to Digital Converter
Report Wave Mixer Error as raw analog output for user reading
Remove XY Motion Control Interface
Extract and report low latency direction information as compass to alignment center based on dither motion
Simplify Digital Signal Processing to simple voting filter to discard false signal and remove 3-point Gaussian Area Algorithm

## Slide 5
System Overview - Block Diagram and Pin count
State Control Mixed Signal System Pin count (14) Analog pin for readout (4) SPI (4) Custom Config, User Command, Status Report Low latency feedback (4) move_en dir ΔΣ Modulated Digital Wave (2) mems_drv

## Slide 6
System Overview - State Machine
Boot
Load Config
Calibration (Optional)
Fallout (Optional)
Readout
Change Amp Ratio
After Load Config Split to 1. Calibration (missing phase_offset from SPI import) Enter Fallout after calibration Reboot system and load phase_offset data
2. Readout

## Slide 7
System Overview - RTL Traffic
Readout
Dynamic Comparator
Wave controller 1. Wave generation (Driver/Reference Wave) - Sigma Delta Modulated 2. Calibration (Optional) 3. Latch pulse (90°/270°)
Signal processor Output move_en, dir
Intermediate RTL for SPI and handshake High: 3.3V Low: 0V Minimize Digital-Analog Crossing to avoid switching noise and complexity

## Slide 8
System Overview - Signal Processor
Comparator set to baseline 1.65V midrail as dc_offset = 0 from wave mixer error output.
Comparator goes through rs latch and inverter to store its value, ready for digital domain to retrieve.
Data is accessed and latch every 90°/270° calibrated phase delay.
Data is filtered through 16 sample voting to reject noise and jitter, latch for low latency output
In-phase: High on 0° - 180°, Low on 180° - 360°
Out-of-phase: Low on 0° - 180°, High on 180° - 360° 90° and 270° are the most stable point for latching and parsing comparator output

## Slide 9
RTL Traffic - Analog Domain
1. Readout
(180μm x 100μm)
2. Wave Mixer (X/Y)
(39.30μm x 55.99μm)
3. Comparator (X/Y)
(67.39μm x 53.38μm)

## Slide 10
RTL Traffic - Analog Front End
Schematic Completed
Layout still WiP
Reserved Area
(180μm x 100μm)
Readout input (South Center)
Transimpedance Amplifier
Band Pass Filter
Readout output (South Center)
Split to Wave Mixer (X/Y)
In symmetry

## Slide 11: Analog Front End - Transimpedance Amplifier
Active Low Enable signal disables current mirror and rest of amplifier, to be controlled by digital signal until config is complete,

## Slide 12: Analog Front End - Transimpedance Amplifier
Gain
DC photodetector current sweep
1uA-40uA linear response
Open Loop 1 Hz ~ 10MHz Well covering target specs
Transient Simulation @156kHz
IIN: 40uA DC, 4uA amplitude
VOUT = 1.47V DC, 0.11V Vpp

## Slide 13: Analog Front End - Two Integrator Loop Gm-C BP + LP Filter
Designed for 156 kHz Tunable for each allowed MEMS frequency 2 Hz ~ 156kHz

## Slide 14: Analog Front End - Testbench of TIA + Bandpass Filter

## Slide 15: Analog Front End - TIA + Bandpass Testbench
Photodetector Sample Input:
40uA DC
4u amplitude
156kHz frequency vs 1MHz shows filtering, though output swing could be improved

## Slide 16
Analog Front End - Full Simulation
Band Pass Filter
X: 300Hz Y: 400Hz
Readout Oscillation
Band Pass Filter
TIA

## Slide 17
RTL Traffic - Wave Mixer & Comparator
Gilbert Cell → 2x 5T OTA → RC Filter → StrongArm Comparator → Inverter x2 → RS Latch → 2x Inverter

## Slide 19: Mixer Schematic
Standard gilbert cell mixer to the left and 5t ota to the bottom
5t ota buffer is needed to convert to single ended and add some gain
5t OTA, current mirror attached to i_bias

## Slide 20: Mixer Testbench
Resulting DC Output to Comparator for algorithm
312kHz
156kHz
DC

## Slide 21
Mixer Testbench with low pass filter (fc = 302.80kHz)

## Slide 22: Comparator Schematic
Simple Strong Arm Latch Comparator
Simple RS Latch for memory

## Slide 23: Comparator Testbench
In_p: wave mixer output In_n: dc offset (V mid rail) Comp_out: digital interface

## Slide 24: Comparator Layout DRC
Can ignore density rules for all metals and poly and active layers for DRC

## Slide 25: Mixer Layout DRC
Can ignore density rules for all metals and poly and active layers for DRC

## Slide 26: Digital - Summary
System Verlog Completed: SPI (100%) State Machine (100%) Wave Controller (100%) Signal Processor (100%)
Cocotb Tested: SPI (90%) State Machine (100%) Wave Controller (100%) Signal Processor (100%)
Gtkwave Timing Visualization: SPI (100%) State Machine (100%) Wave Controller (100%) Signal Processor (100%)
Underdevelopment: Amp Ratio Adjuster
Links: 1. System Verilog HDL Probton\src https://github.com/llhtimlam/Probton/tree/main/src
2. Cocotb Simulation Probton\cocotb https://github.com/llhtimlam/Probton/tree/main/cocotb
3. Gtkwave Probton\cocotb\sim_build
https://github.com/llhtimlam/Probton/tree/main/cocotb/sim_build

## Slide 27: Digital - State Machine Simulation
Boot -> Load Config -> Calibration -> Fallout (cal_done or cal_timeout) Boot -> Load Config -> Readout -> Change Amp Ratio -> Readout Passed test on all transition

## Slide 28: Digital - Wave Controller
Sin wave is generated for memes driver (expecting triangular wave but use sin wave due to third odd harmonic) Square wave is generated for wave mixer reference wave Utilized Sigma Delta Modulator for pwm wave generation

## Slide 29
Calibration: Phase-Lag Capture with Jitter Rejection
Jitter guard active
Early/invalid edges rejected (wave_is_valid=0)
Capture re-arms, no offset committed
Valid Calibration
All three edges captured (wave_is_valid=1)
90°/270° offsets computed: ~523k / ~1573k in phase-accumulator units

## Slide 30
Readout: Phase-Sampling Strobes with Ack Handshake
90° strobe
Fires at phase_acc = 524288 (2²¹/4), held until ack, then cleared
270° strobe
Fires at phase_acc = 1572864 (3·2²¹/4), held until ack, then cleared

## Slide 31
Simulation Result
Improvement could be made by voting filter and adjusting amp ratio for better resolution
Just using the direction bit retrieved from the wave mixer, it is able to locate the center but converge too early.

## Slide 32
Simulation Result
Improvement could be made by voting filter and adjusting amp ratio for better resolution
Shows concentrated convergence result due to early converges and not able to identify the direction near peak

## Slide 33
Project Planning and Tracker

## Slide 34
Question
1. Should we bundle readout intensity with wave mixer error for synchronous digital reporting, or maintain isolated asynchronous paths in analog?
2. How should we change the amp ratio? Should we use a switch to disable the readout and then change the analog setting?
3. Feedback on the chip architect and whole system?
Risk (Looking for advice)
1. Since we decided to skip PLL and use sigma delta modulated for digital wave generation using 8-bit sin LUT, there will be some issue with harmonic. Design solution with eternal RC or alternative wave mixer component that compatible with digital reference wave are considering.
2. Amp Ratio Adjuster might be tricky as it need a lot of wires from digital side for switching resistor feedback.
2. Since we only consider 1 comparator park at 1.65V midrail, it might suffer jittering during idle at 0 offset, which might skew the result from calibration. Hystereis design might be required.

## Slide 35
References
[1]. H.-S. Liao *et al.*, "Low-cost, open-source XYZ nanopositioner for high-precision analytical applications," *HardwareX*, vol. 11, p. e00317, 2022. doi: [10.1016/j.ohx.2022.e00317](https://doi.org/10.1016/j.ohx.2022.e00317)
[2]. W.-M. Wang *et al.*, "Low-voltage and high-performance buzzer-scanner based streamlined atomic force microscope system," *Nanotechnology*, vol. 24, no. 45, p. 455503, Nov. 2013. doi: [10.1088/0957-4484/24/45/455503](https://doi.org/10.1088/0957-4484/24/45/455503)
[3]. Rout, Shasanka Sekhar, et al. "Design of 2.4 GHz Improved Current Reuse Gilbert Mixer with Source Degeneration Technique." Wireless Personal Communications, vol. 122, 2022, pp. 3875–3887, https://doi.org/10.1007/s11277-021-09115-6.
[4]. Altuner, Emre, et al. "High-Linearity Gilbert-Cell Mixer Design for Cryogenic Applications." Analog Integrated Circuits and Signal Processing, vol. 113, 2022, pp. 249–256, https://doi.org/10.1007/s10470-022-02098-9.

## Slide 36: Team
Tim (Team Lead)
HKUST Chem (UG)
UW HackerFab UWASIC
Abraar TMU CE (UG) LCRAIN
Reza UCD ECE (PG)
Victoria UW NE (UG)
UWASIC
IQC UG RA
Ashmita UCD ECE (PG)
Belkacem UT EE (PhD) Automation Lab
Nitin UCSB EE (PG) Mediatek AMS Design engineer
Annika
Kalam HKUST EE (PG)